## Slide 1: Cochlear implant telemedicine to overcome geographic barriers and increase access to care Joshua Bernstein PhD1, Elicia Pillion AuD1, Sandeep Phatak PhD1, LTC Anthony Tolisano MD2 1Audiology & Speech Pathology Center, 2Department of Otolaryngology—Head and Neck Surgery
Cochlear implants (CIs) are a spectacularly successful neural-prosthetic treatment for severe to profound hearing loss.
Most active-duty service members who receive a CI are able to remain on active duty.
Geographical restrictions are a major barrier to access to CI due to limited number of MTFs that support care.
CI recipients require extensive travel and mission disruption, or civilian network care costing the government >$35k per patient.
This study developed a prototype telemedicine approach, allowing patients to complete nearly all their follow-up care without traveling to the CI center.
Two studies evaluated validity (comparison to standard of care measures) and efficacy (field trial evaluation).
Calibration. Headphones placed over the CI sound processor. Virtual audio techniques, using acoustic manikin and “cochlear implant in a box”, to determine headphones filter needed to produce the same electrical response as loudspeaker testing
Study 1: Validation. Compare telehealth system headphones-based speech perception and pure-tone thresholds to standard-of-care loudspeaker sound-field testing (N=15)
Study 2: Field study. Complete CI audiological follow-up appointments from the patient’s home, with patients completing questionnaires to rate their experience (N=13)
INTRODUCTION
METHODS
RESULTS
CONCLUSIONS
 | Standard practice | Telemedicine
Audiogram | Sound field | Headphones
Speech tests | Sound field | Headphones
Programming | Company software USB connection In-person | Company software USB connection Remote control
Sound quality check | Live voice | Pre-recorded sounds, headphones
Questionnaires | Pencil & paper | Online
Red Font = Different from standard practice
FIGURE 1: Standard and telehealth systems produce similar CI “sound-field” audiograms
(a) Group-mean audiograms are indistinguishable for the two systems
(b) Loudpeaker-headphone differences are within expected test-retest reliability limits: 95% confidence interval (95ci) and limit of agreement (LOA)
FIGURE 2: Standard and telehealth systems produce similar speech-testing results
All elements of a post-CI visit replicated in the telehealth system
Telehealth headphones-based speech scores are within or close to the 95% confidence interval. (a) consonant-nucleus-consonant (CNC) phonemes, (b) CNC words, (c) sentences in quiet, (d) sentences in noise
FIGURE 3: Field study shows high patient satisfaction
Subjects completed a survey following each telehealth visit. 39 responses from 13 subjects show high patient satisfaction with telehealth.
Mean absolute differences between telehealth-system and sound-field audiograms were 5 dB. 96% of 500–4000 Hz telehealth-system thresholds within the sound-field 95ci.
CNC word and sentence identification in quiet showed high interclass correlations (ICC=0.96–0.98) and 80-93% of telehealth-system scores within the sound-field 95ci. Sentences in noise scores showed larger deviations.
13 CI patients located across the U.S. and overseas completed 39 post-op telehealth programming and assessment visits.
Telehealth-survey responses indicated high patient satisfaction, improved convenience and access to care, and only minor challenges.
The headphone-based system for at-home follow-up appointments is accurate, feasible, and satisfactory to patients.
By reducing geographical barriers and increasing access to care, the telehealth approach has the potential to
  increase return-to-duty for hearing-impaired service
  limit mission impact by maintaining continuity of care regardless of geographical requirements
  recapture patients from the civilian network, saving the government 10s of thousands of dollars per surgery.